| Instrumentation programme | Report 04.2 |
| --- | --- |

RESULTS · THERMAL DRIFT

# Drift stops being correctable above 50 °C

Below 50 °C the error is linear and compensates cleanly in software. Above it the response becomes hysteretic, and the same correction no longer holds.

TABLE 1 · DRIFT BY AMBIENT BAND

| AMBIENT | DRIFT | RESPONSE | N |
| --- | --- | --- | --- |
| 20 °C | 12 ppm | Linear | 12 |
| 35 °C | 34 ppm | Linear | 12 |
| 50 °C | 88 ppm | Mixed | 12 |
| 65 °C | 310 ppm | Hysteretic | 10 |

± 1σ across units · n=46 · illustrative data · thermal chamber, 2026

DRIFT INCREASE, 20 → 65 °C

26×

An order of magnitude, and a change of kind.

## Method

Forty-six units measured against reference across four ambient bands. Uncertainty is stated at one sigma throughout. Where a correction is quoted, the region of validity is quoted with it.

## What this changes

A compensation built on the linear region cannot be extended past it. Either the operating envelope is capped at 50 °C, or the correction is rebuilt on a model that admits hysteresis.

SCHEDULE

| ACTION | OWNER | DUE |
| --- | --- | --- |
| Cap operating envelope | Firmware | 14 Nov |
| Rebuild correction model | Metrology | 12 Dec |
| Re-test at 65 °C | Metrology | 09 Jan |